Thermo Scientific Barnstead Megapure Still Pre-Start-Up Checklist

| Instructions for Use: Please complete all sections of this Site Inspection form in preparation for the start-up of your new water purification system. This is a fillable form. Failure to have the start-up site ready will result in delays and additional charges. See FAQ for water system start-up at www.thermofisher.com/installation Rename document with your company name at the end. Save completed form for your own records. Email the completed form to: servicesupport.led.asheville@thermofisher.com |
| --- |
| Site Name |  |
| Address |  |
| Address/Bldg. Room Number |  |
| City |  | State |  | Zip |  |
| Lab Hours (from-to) |  | Weekday |  |
| Lab Contact Name |  | Lab Contact Phone | Extension ( ) - Extension |
| Lab Contact Email |  |
| Alternate Contact |  | Alt. Contact Phone | Extension ( ) - Extension |
| Alt. Contact Email |  |
| Purchase Order |  | Case number (from email notification) |  |
| Special Instructions to get into facility |  |

Section 1 - System Identification

Section 2 - Accessories and Unpacking

Section 3 - Location Requirements

Section 4 - Electrical Requirements

Section 5 - Water and Drain Specifications

Section 6 - Other Specifications and Additional Notes

Section 7 - Authorized Certification & Customer Agreement

Section 8 - Thermo Scientific Review

Section 9 - Addendum

Section 2:  Accessories and Un

| Section 1: System Identification Check the Model and enter the Serial Number for the system. If multiple units of system are to be set-up, please list all serial numbers below. |
| --- |
| A440266 MP-1 120 V A440267 MP-1 120 V with 6 L bottle A440367 MP-3A 240 V A440696 MP-3A 208 V A440518 MP-6A 240V A440697 MP-6A 208V |
| System Serial Number: |  |
| Is this unit replacing an existing water purification system? If Yes, what is the Model number of the unit being replaced? Serial number? If No, skip to section 2. |  |
| Note: Existing system must be removed and disposed of by customer |
| Section 2: Accessories Indicate which accessories were ordered/received |
| Feed Water Pretreatment Accessories - Optional: 440236 Dual Solenoid Valve (MP-3A and MP-6A) RY798X2A Dual Solenoid Valve (MP-1) D4511 Single B-pure 331040 Adapter Kit for D4511 Qty 1 D0803 High capacity cartridge D4524 Double B-pure 331041 Adapter Kit for D4524 Qty 2 D0803 High capacity cartridge 440092 Flow Meter (MP-1 with use with Bpure pretreatment) Storage - Required 413964 6 Liter Plastic Bottle 413934 13 L Glass Bottle B440704 Automatic Collection System (ACS) ACS Accessories- Optional: 440138 Flexible Tube Adapter Kit (Customer must supply 0.5” Teflon tubing) 400634 ACS Wall Bracket 440241 Adapter for ACS washer hook up |
| Section 2: Unpacking Unpack system and inspect glass for any shipping damage. Note, system and accessories will arrive in separate packages. Locate accessories box and heating elements. If any components appear to be missing or damaged, please contact. customerservice.led.asheville@thermofisher.com . Include model, serial number of system and missing/damaged component. |
| System was unpacked and contents was reviewed | Yes No, unit is still in box |
| Section 3: Location Requirements It is the responsibility of the customer’s onsite facilities department to: Review space requirements, mounting instructions and precautions/warnings per system operation manual. See Addendum for general layout configurations. Set up water system in its final location. Wall Mounting: Customer is responsible for mounting MEGA-Pure Still and Automatic Collection System wall bracket to the wall prior to start-up. MEGA-PURE Water Still has slotted holes for wall mounting. To support the weight of this unit, customer must supply and install 1/4" x 1-1/4" lag bolts for mounting. An optional wall mounting bracket is available for the Automatic Collection System. Thermo Scientific Bpure pretreatment must be wall mounted only. Wall bracket is included. Screws and fasteners required for wall mounting are not supplied with the unit. Bpure pretreatment location is not critical; however, it should be close enough to the still to make all necessary connections. Bench Mounting: If Bench Mounting both MegaPure Still and ACS, the MegaPure Still must be mounted to a customer provided, secure, 12" platform to allow for gravity flow of distilled water to the ACS. Drawings with dimensions are in the operation manual. It is the responsibility of Unity Lab Service to do the following: Set up equipment with tubing and accessories. Install cartridges and filters. Startup system and check for leaks and proper operation. Training on operation and maintenance of system. Start-up service does not include drilling into walls or countertops. Please complete the section below to verify where the system and accessories will be mounted. Indicate below if the system and accessories have been mounted in their final location. |
| Is MegaPure system mounted in final location? | Yes, wall mounted Yes, bench mounted on stand No, not ready * |
| Is Automatic Collection System mounted in final location? | Yes, wall mounted Yes, bench mounted on stand No, not ready * NA |
| Is Bpure Pretreatment system mounted in final location? | Yes, wall mounted No, not ready * NA |
| * If marked “Not Ready”, the water system must be in location or wall mounted before the service engineer arrives. | Date System will be ready: (MM/DD/YYYY) |
| Section 4: Electrical Requirements The water still is supplied with a power cord and plug. Customer must provide a dedicated power supply to match system power requirements: 120V (MP-1), 240 VAC or 208 VAC (MP-3A or MP-6A) within 3-5 ft. of Megapure unit. Refer to the “Electrical Requirements” section of the operation manual for the MegaPure Glass Still for power requirements and plug NEMA configuration. |
| Have Electrical requirements been met? | Yes No |
| Section 5: Feed Water & Drain Requirements The MegaPure water system comes with a solenoid valve with 1/8” NPTF connection. See Section 9 Addendum for picture of supplied adapter. It is the responsibility of the customer’s onsite facilities department to: Customer must supply shutoff valve with 1/8" NPT reducer within 3-5 ft. of Megapure still. See Dual Solenoid valve accessory: If system is fed with Tap water and Pretreated source, shut off valve and with 1/8" NPT reducer is required for both water sources. If water supply does not have correct connector, delays in start-up and additional charges will be incurred. It is recommended to include a photo of the feed water connection. |
| Feed water has shutoff valve with 1/8" NPT reducer within 4 ft. of Megapure still? | Yes No* |
| *If no was selected, please describe water connection in additional notes section below and/or supply photo. |
| Feed water temperature between 4 - 37 °C? | Yes No |
| Feed water pressure between 1.4 - 6.2 bar (20 - 90psi)? | Yes No |
| Atmospherically vented drain or sink within 3-5ft, can handle up to 76 LPH (20 GPH) flow and 200°F and it is at a downward slope from still? | Yes No |
| Section 6: Other Specifications and Additional Notes The primary operator of the instrument must be available during start-up visit and for training during the initial instrument setup. |
| Will someone be available? | Yes No |
| ADDITIONAL NOTES: Please explain any “No” or “Not Ready” responses in Sections 1 – 6 above. Also enter additional notes and comments about the installation site or system set up in the section below. For example, is the Megapure Still going to be connected to another system? |
| Section 7: Customer Agreement I have reviewed the attached site inspection form and agree that all specifications required for the start-up of the Water Purification system can be met unless otherwise noted on this form. Customer acknowledges it is their responsibility to complete the activities listed above. If the site is not ready when the technician arrives and the start-up cannot be completed, the customer will be responsible for charges including time and travel associated with a second service call to complete the start-up, minimum fee of $300 to be charged. Incorrectly filled out forms or misrepresentation may result in additional charges at the discretion of Thermo Fisher Scientific or their authorized partner. Please consult with your service engineer or sales representative with any questions. |  |
| Customer name Print Name |  |
| Customer Signature Signature |  |
| Date |  |
| Section 8: To Be Completed By Authorized Thermo Fisher Service Representative On-Site I have reviewed the attached, completed customer site inspection form and confirm that the specifications for the start-up were met. |
| Name of Thermo Scientific Service Representative |  | Yes No* |
| *If No, please make a note of missing or incorrect specifications and forward a copy of the completed form to Service Manager and Sales Representative for consultation: |
| Notes and Comments: |

Section 9 – Addendum

Feed water supply connection supplied by customer requires:

- Water supply shut off valve

- At least 2-3” of pipe /tubing

- 1/8” NPT male connector to connect to Megapure feed water solenoid valve.

Example of Megapure with fed with tap water:

Example of Megapure with deionizer (Bpure) fed water:

Example of Megapure with tap and house treated water:

Example layout of Megapure Still, Bpure Pretreatment and ACS