Thermo Scientific Barnstead Pacific TII or RO Pre-Start-Up Checklist

| Instructions for Use: Please complete all sections of this Site Inspection form in preparation for the start-up of your new water purification system. This is a fillable form. Failure to have the start-up site ready will result in delays and additional charges See FAQ for water system start-up at www.thermofisher.com/installation Rename the document with your company name at the end. Save completed form for your own records. Email the completed form to: servicesupport.led.asheville@thermofisher.com |
| --- |
| Site Name |  |
| Address |  |
| Address/Bldg. Room Number |  |
| City |  | State |  | Zip |  |
| Lab Hours (from-to) |  | Weekday |  |
| Lab Contact Name |  | Lab Contact Phone | Extension |
| Lab Contact Email |  |
| Alternate Contact |  | Alt. Contact Phone | Extension |
| Alt. Contact Email |  |
| Purchase Order |  | Case number (from email notification) |  |
| Special Instructions to get into facility |  |

Section 1 - System Identification

Section 2 - Accessories and Unpacking

Section 3 - Location Requirements

Section 4 - Electrical Requirements

Section 5 - Water and Drain Specifications

Section 6 - Other Specifications and Additional Notes

Section 7 - Authorized Certification & Customer Agreement

Section 8 - Thermo Scientific Review

Section 9 -  Addendum

Section 2:  Accessories and Un

| Section 1: System Identification Check the Model and enter the Serial Number for the system. If multiple units of system need start-up, please list all serial numbers below. |
| --- |
| Pacific TII Systems: 50132123 – Pacific TII 7 50132131 – Pacific TII 7 UV 50132124 – Pacific TII 12 50132132 – Pacific TII 12 UV 50132125 – Pacific TII 20 50131982 – Pacific TII 20 UV Pacific RO Systems: 50132386 – Pacific RO 7 50132387 – Pacific RO 12 50132388 – Pacific RO 20 |
| System Serial Number: |  |
| Is this unit replacing an existing water purification system? If Yes, what is the Model number of the unit being replaced? Serial number? If No, skip to section 2. |  |
| Note: Existing system must be removed and disposed of by customer |

| Section 2: Accessories – Pretreatment Required Indicate which Pretreatment was ordered/received. Indicate if Storage tank and wall bracket was ordered/received. |
| --- |
| 50157886 - Barnstead Pretreatment System 09.4000 - Dual Pretreatment Pretreatment Not purchased* *Please consult your sales representative or service engineer for guidance. |
| 50138221 Hand Dispenser – Optional Accessory |
| Section 2: Accessories – Storage Tank Required Indicate which storage tank, wall bracket and tank accessories were ordered/received. |
| Tanks 50155495 - 30 L tank 50155494 - 30 L tank w/pump 50155500 – 60 L tank 50155499 – 60 L tank w/pump 50155503 – 100 L tank 50155502 – 100 L tank w/pump Tank not purchased/received* *Please consult your sales representative or service engineer for guidance. Required Tank Accessories: 06.5001 – Sterile Overflow 06.5002 – CO2 Vent filter 50135142 Vent filter Optional Tank Accessories: 06.5015 30 L Tank Wall Bracket 065016 60 L Tank Wall Bracket 50155463 2 -Outlet Manifold 50155458 3 - Outlet Manifold |
| Section 2: Unpacking Unpack system. Note, systems will arrive in multiple boxes. Locate accessory parts bags and DI Cartridge (Pacific TII only) and review contents as listed in “Extent of Assembly Kit” section found in water system operation manual. If any components are missing, please contact customerservice.led.asheville@thermofisher.com . Include model, serial number of system and missing component. |
| System was unpacked and “Extent of Assembly Kit” was reviewed | Yes No, unit is still in box |
| Section 3: Location Requirements It is the responsibility of the customer’s onsite facilities department to: Review space requirements, mounting instructions and precautions/warnings per system operation manual. Ensure the bench or wall can support the weight of the instrument per the Technical Specifications in the operation manual: Pretreatment = 50 lbs., Pacific system = 60 lbs., 30 L tank = 90 lbs., 60 L tank = 160 lbs., 100 L tank = 254 lbs. If wall mounting, wall mounting hardware MUST be attached to the wall prior to the date of start-up. The system or wall bracket comes with the required mounting hardware. Set pretreatment, water system, and tank in its final location. Drawings available at www.thermofisher.com/newlab It is the responsibility of Unity Lab Service to do the following: Set up equipment with tubing and accessories. Install cartridges and filters. Startup system and check for leaks and proper operation. Training on operation and maintenance of system. Start-up service does not include drilling into walls or countertops. Please complete the section below to verify where the system and accessories will be mounted. Indicate below if the system and accessories have been mounted in their final location. |
| Where Pretreatment system will be mounted? 09.4000 wall mount only 50157886 wall or bench mount options -wall bracket included | Wall Mount Bench Mount |
| Is Pretreatment noted above mounted in final location? 3- 6 ft. from Pacific system | Yes No, not ready * |
| Where Pacific system will be mounted? Wall bracket ordered separately with exception of Pacific System Package with Tank and pretreatment | Wall Mount Bench Mount Undercounter Mount |
| Is Pacific system mounted in final location? | Yes No, not ready * |
| Is storage tank mounted in final location? Maximum distance of 6 ft. from Pacific system and have gravity flow access to a drain or sink. | Yes No, not ready * |
| * If marked “Not Ready”, the water system must be in location or wall mounted before the service engineer arrives. | Date System will be ready: (MM/DD/YYYY) |
| Section 4: Electrical Requirements System is supplied with a power supply and plug options. Customer must provide a grounded 100 – 250 VAC, 50 – 60 Hz, 2.0 A max electrical outlet within 5 ft. of the installation location. See Electrical Requirements section in the operation manual. |
| Have Electrical requirements been met? | Yes No |
| Section 5: Feed Water & Drain Requirements – 09.4000 pretreatment has ¾” NPTF feed water adapter and 6 ft of feed water tubing. 50157886 Pretreatment has feed water adapter options, ¼” NPTM or ¾” NPTF, and 6 ft of tubing. See Section 9 Addendum for picture of supplied adapter. The Pacific water system comes with ¾” NPTF adapter to connect to pretreatment as well as tubing for feedwater, drain and for storage tank connections. It is the responsibility of the customer’s onsite facilities department to: Feed water supply must be tap water. Customer must supply ¼” NPTF or ¾” NPTM at the water source to connect to the water purification system feed water adapter. If water supply does not have correct connector, delays in start-up and additional charges will be incurred. It is recommended to include a photo of the feed water connection. Check FAQ at www.thermofisher.com/installation on how to identify per feed water connection. |
| Feed water supply fitted with ¾” NPT (pipe thread) male connector or ¼” NPT connector? | Yes No* |
| *If no was selected, please describe water connection in additional notes section below and/or supply photo. |  |
| Feed water has shutoff valve? | Yes No |
| Feed water supply located within 3 - 6 ft. of pretreatment? | Yes No |
| Feed water temperature between 2-35°C? | Yes No |
| Feed water pressure between 2-6 bar (29 – 87psi)? | Yes No |
| Atmospherically vented drain for RO permeate and storage tank overflow within 5 ft. of system and at a downward slope? | Yes No |
| Section 6: Other Specifications The primary operator of the instrument must be available during start-up visit and for training during the initial instrument setup. |
| Will someone be available? | Yes No |
| ADDITIONAL NOTES: Please explain any “No” or “Not Ready” responses in Sections 1 - 6 above. Also enter additional notes and comments about the start-up site or system set up in the section below. For example, is the Pacific storage tank feeding another system? |
| Section 7: Customer Agreement I have reviewed the attached site inspection form and agree that all specifications required for the start-up of the Water Purification system can be met unless otherwise noted on this form. Customer acknowledges it is their responsibility to complete the activities listed above. If the site is not ready when the technician arrives and the start-up cannot be completed, the customer will be responsible for charges including time and travel associated with a second service call to complete the start-up (minimum time/travel fee of $300). Incorrectly filled out forms or misrepresentation may result in additional charges at the discretion of Thermo Fisher Scientific or their authorized partner. Please consult with your Field Service Engineer or Territory Sales Representative with any questions. |
| Customer name Print Name |  |
| Customer Signature Signature |  |
| Date |  |
| Section 8: To Be Completed By Authorized Thermo Fisher Service Representative On-Site I have reviewed the attached, completed customer site inspection form and confirm that the specifications for start-up were met. |
| Name of Thermo Scientific Service Representative |  | Yes No* |
| *If No, please make a note of missing or incorrect specifications and forward a copy of the completed form to Service Manager and Sales Representative for consultation: |
| Notes and Comments: |

Section 9:  Addendum