{"text_format":"1","workbook":"xlsx","n_sheets":1,"sheets":["IQ_Report"],"not_rendered":{"charts":0,"images":1,"macros":1,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"IQ_Report","index":0,"n_rows":46,"n_cols":9,"rows_kept":46,"cut":false,"hidden":false,"merged":[[45,2,45,3],[44,2,44,3],[37,1,37,4],[5,7,6,8],[30,8,32,8],[25,8,27,8],[0,0,1,5],[31,1,31,4],[20,1,21,1],[20,2,21,2],[2,1,2,4],[3,1,3,4],[6,1,6,4],[8,1,8,4],[18,1,18,4]]}
[]
[]
[null,"Installation Qualification"]
[null,"Vanquish Neo System"]
[]
[null,null,null,null,null,null,null,"NOTES"]
[null,"Retention Time Reproducibility and PWHM in MS or UV"]
[]
[null,"Configuration",null,null,null,null,null,null," = are required fields to be filled out."]
[]
[null,"Instrument Component","P/N / Version","  Supplier's name","  Serial number",null,null,"Configuration","If particular instrument components are not used / installed, please fill in \"n.a.\"."]
[null,"Autosampler",null,"Thermo Scientific",null,null,null,null,"For the P/N, use the currently installed module / software version"]
[null,"Pump module",null,"Thermo Scientific"]
[null,"Column oven",null,"Thermo Scientific",null,null,null,"Calculation","Use the last 6 runs of the IQ sequence to calculate the retention time reproducibility"]
[null,"Vanquish VWD F/C",null,"Thermo Scientific"]
[null,"SII version",null,"Thermo Scientific"]
[null,"MS software",null,null,null,null,null,"Test Results","Fill out the test results according to the detector used. "]
[null,null,null,null,null,null,null,null,"For an LC-MS configuration, use the top two criteria to determine the IQ results"]
[null,"Calculation retention time reproducibility",null,null,null,null,null,null,"For an LC-UV configuration use the top and bottom criteria to determine the IQ results"]
[]
[null,"Run no","Name ","Retention time (min)",null,null,null,"Date","Verify the electonic assignment of the date matches the instalaltion date. "]
[null,null,null,"m/z 728.84 (2+)",null,null,null,null,"Manually correct this if necessary"]
[null,1,"Cytochrome C Digest"]
[null,2,"Cytochrome C Digest"]
[null,3,"Cytochrome C Digest"]
[null,4,"Cytochrome C Digest",null,null,null,null,null,"Click this button to clear all entered values in the IQ Report."]
[null,5,"Cytochrome C Digest"]
[null,6,"Cytochrome C Digest"]
[null," Average:",null,{"$f":"IF(COUNT(D23:D28)=6,AVERAGE(D23:D28),\"Enter RT values \")","v":"Enter RT values "},"Min"]
[null," RSD:",null,{"$f":"IF(ISNUMBER(D29),(STDEV(D23:D28)/AVERAGE(D23:D28))*100,\"Enter RT values \")","v":"Enter RT values "},"%"]
[null,null,null,null,null,null,null,null,"Click this button to print the IQ report as PDF. \nEither print it in paper for signing or use electronic signatures"]
[null,"Test results LC-MS / LC-UV configuration"]
[]
[null,"Test","Limit","Observed values","Result"]
[null,"Retention time repro. ","≤ 0.30% RSD",{"$f":"D30","v":"Enter RT values "},{"$f":"IF(ISNUMBER(D35),IF(D35>=0,IF(D35<=0.3,\"Passed\",\"Failed\"),\"Enter correct value\"),\"Enter correct value\")","v":"Enter correct value"}]
[null,"PWHM  (s) ","≤ 7 s",null,{"$f":"IF(ISNUMBER(D36),IF(D36>=0,IF(D36<=7,\"Passed\",\"Failed\"),\"Enter correct value\"),\"Enter correct value\")","v":"Enter correct value"}]
[]
[null,"Additional results LC-UV configuration"]
[]
[null,"Test","Limit","Observed values","Result"]
[null,"Resolution Peak 1-4","≥ 1.5",null,{"$f":"IF(ISNUMBER(D41),IF(D41>=0,IF(D41>=1.5,\"Passed\",\"Failed\"),\"Enter correct value\"),\"Enter correct value\")","v":"Enter correct value"}]
[null,"Peak 5 Height [mAU]","≥ 3",null,{"$f":"IF(ISNUMBER(D42),IF(D42>=0,IF(D42>=3,\"Passed\",\"Failed\"),\"Enter correct value\"),\"Enter correct value\")","v":"Enter correct value"}]
[]
[]
[null,null,null,null,{"$f":"TODAY()","v":"2022-11-21"}]
[null,"Customer's signature","Operator's signature",null,"Date"]