## TMO 5.5 HP Troubleshooting KML
Symptom | Potential Cause | Diagnostic File | What to Inspect | Resolution / Service Manual | Parts / Kits Needed for repair | Tools | Repair validation | Replace UNIT under warranty Y/N
Fan Error
Unit cannot reach or hold its set point temperature | Ensure both external fans operating properly. |  | Ensure Fan1 / Fan 2 are not blocked or not spinning due to debris. Inspect fans for physical damage indicating blockage occurred. | Inspect front fan grates for blockage; Remove blockage if present, and ensure adequate ventilation space behind the product. | Fan kit | No tools needed | Power on, ensure both fans are operating for 5 minutes. Reassemble unit. | No
|  |  | Ensure Fan1 / Fan2 connector is not disconnected. | Inspect connection at power supply; remove wiring connector and inspect pins on both sides; reseat connector. If pins on Power Supply are damaged beyond repair - replace power supply. | Fan kit &/or Power Supply Kit | Precision screwdriver, Precision tweezers. *Mobile camera to take picture of connector to inspect for pin damage. | Power on, ensure both fans are operating for 5 minutes. Reassemble unit. | Yes due to cost and lead time of P/S kit
|  |  | Ensure Fan1 / Fan2 cable is not damaged. | Inspect wiring cable between power supply and Fan1 or Fan2 for damage. Inspect for loose wires in the connector. | Fan kit | No tools needed | Power on, ensure both fans are operating for 5 minutes. Reassemble unit. | No
0x0008 air and 0x0010 indicate a control senor issue
Unit cannot reach or hold its set point temperature | User Interface Display Accuracy |  | Check cabinet temperature with an independent thermometer | If cabinet is at proper temperature, replace the air sensor | Temperature sensors kit | Heat shrink tub, solder, wire strippers, wire crimpers | Monitor pulldown to ensure display temperature is accurate with external thermometer. | Need to test for time involved
|  |  | Ensure the air sensors are not damaged or disconnected from power supply. | Inspect air sensor element or cable assembly for damage. Remove air sensor cable assembly for damage; Disconnect air sensor harness from power supply. Reseat connector to power supply. | Temperature sensors kit | Calibrated Fluke (or similar) thermometer, Phillips Screwdriver, Needle nose pliers (for air sensor), | Confirm critical system error clears and temperature is accurate with external thermometer. | Or if a soldering iron is required
0x0080 indicates a critical reject temperature.
Unit cannot reach or hold its set point temperature | Confirm Reject Temperature hasn't reached critical limit | AmbientTemp > 32 | Inspect environment to ensure ambient temperature is not above 32C. | Install in a proper environment | N/A | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  | N/A | Feel for air flow on the front right of the unit. | Inspect front fan grates for blockage; Remove blockage if present, and ensure adequate ventilation space in front of the product. Ensure front reject is clear from obstructions | N/A | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  | N/A | Ensure Heat reject fins are not blocked, clogged, or heavy dust/particulate | Clean unit to remove dust & debris | Fan kit | Phillips Screwdriver, Vacuum/Compressed Air can for dusting. | Power on and inspect fans for operation. | No
|  |  | Ensure both fans are spinning. | See section on fan functionality | N/A | N/A | N/A | No
Critical power supply issues to the PHP
Unit cannot reach or hold its set point temperature | Ensure PHP is working properly |  | Check PS voltage to PHP | See Service guide for Wiring diagram and voltage from Power supply to components | Power Supply Kit | Volt Meter, Precision tweezers, Phillips Screwdriver | Is voltage output correct? PHP failure is possible cause. | Yes due to cost and lead time
|  |  | Remove the top cover, inspect the PHP mounting brackets to ensure properly tension on screws and system assembly | If loose then inspect for PHP mounting bracket damage. Tighten screws to 20 in/lbs. | Torque each nut to 20 inch pounds ± 2 inch pounds. Alternate tightening down nuts with incrementally increase in torque setting (start at ~10 inch pounds go to 20 pounds) to determine if nuts were loose before reaching 20 inch pounds setting. CAUTION: DO NOT OVERTIGHTEN, this will damage PHP and void warranty. | 8mm Socket, Torque driver(Utica 10-30 inch drivers, TS-30), Phillips Head Screwdriver | Monitor pulldown temperature based on pulldown chart | Would not expect the field techs to have the tools for this
|  |  |  | If all other tests indicate faulty PHP | PHP Replacement Kit | N/A | N/A | Yes for now
|  |  | Check for a solid green LED light lit on the power supply when unit is plugged in. | Power on unit looking for green LED light on power supply. If no green light, inspect cable from power receptacle to power supply, reseat cable and reapply AC power. If problem persists, replace power supply. | Power Supply Kit | Phillips Screwdriver | Power on, ensure both fans are operating for 5 minutes. Reassemble unit. | Yes due to cost and lead time
|  |  | Ensure all cables are well seated into the power supply. Inspect the crimp pins inside the connector housing to ensure that they are not pulled out | Remove then reseat any cable that appears loose or dislodged. If everything is checked and is properly plugged in, replace the Power Supply | Power Supply Kit | Precision tweezers | Power on, ensure unit pulls down to set point temperature | Yes due to cost and lead time
|  |  | Ensure good PHP interface to thermal system by inspecting the PHP mounting bracket to ensure properly tension on screws and system assembly. | See section on PHP functionality | N/A | N/A | N/A | No
Unit does not cool to the set point
Unit cannot reach or hold its set point temperature | Gasket not sealing properly | Air1, Air2, Glycol_Temperature are greater than 1C higher or lower than the Setpoint_temp. | Ensure chamber is empty, door is closed, and door gasket seals effectively | Confirm proper installation for all items, Put flash light inside & look for gaps of light, if light is visible then door seal is compromised. Can you slide paper through the gap without resistance? Inspect gasket for tears. | Door Gasket | Plastic putty knife | Confirm seal looks & feels, repeat test for door seal. | No
| Door not fully closing |  | Ensure door hinges pull door closed | Conduct hinge replacement | Door Hinge Kit | Hex driver and hex socket M8 - 2 Person operation | Confirm seal looks & feels, repeat test for door seal. | No
| Ice has formed blocking air flow |  | Ensure no ice has formed on walls or door thermal break | Thaw and ensure thermal break seals with door gasket | Door Gasket | Plastic putty knife | Confirm seal looks & feels, repeat test for door seal. | No
| Debris lower efficiency of the unit |  | Inspect the heat fins behind the external fans for excessive dust / particulates | Remove external fans and visually inspect for cleanliness | NA | Vacuum, Compressed Air can for dusting. | Power on, ensure both fans are operating for 5 minutes. Reassemble unit. | No
|  |  | Inspect the filter on filter installed units | Inspect for clogged filter |  | NA | Power on, ensure refrigerator is pulling down the temperature to desired level | No
| PHP Performance |  | Inspect PHP wiring & connector | Remove AC power; re-apply AC power after waiting 30 seconds; If symptom reappears, remove AC power, remove top cover and inspect PHP harness for damage or faulty connection to power supply. Remove & replace PHP connector to ensure proper seating. Replace power supply if symptom continually reappears after cycling AC power. | Power Supply Kit or Power Supply Harness | Precision tweezers, Phillips Screwdriver | Power on unit, Monitor pulldown temperature based on pulldown chart | Yes due to lead time
| Confirm CO2 Charge is adequate |  | Inspect for leaks at charging ports. | If leaking at charge port then system needs to be evacuated, pin valve replaced, and recharged with CO2. Warranty return if leak detected | CO2 Charging Kit | Big Blu Micro Leak Detector or Nu-Calgon's Cal-Blue, pin valve charge fitting, pressure gauge, | NOTE: testing CO2 pressure will void warranty since process will discharge enough CO2 to affect system performance. Only certified technician should handle CO2 system | YES
|  |  | Check pressure of CO2 charge | Testing CO2 pressure will If chamber / unit is at ambient temperature (22C) and CO2 charge is below 800 psi, consult supplier for warranty return. | CO2 Charging Kit | Big Blu Micro Leak Detector or Nu-Calgon's Cal-Blue, pin valve charge fitting, pressure gauge, | NOTE: testing CO2 pressure will void warranty since process will discharge enough CO2 to affect system performance. Only certified technician should handle CO2 system | YES
Temperature is more than 1C colder than the set point
Temperature is greater than 1C colder than the set point | Product added is in contact with a temperature sensor. | Confirm that Air1, Air2, and glycol_temperature are not within 1C setpoint_temperature. Identify any sensors our of range. | Inspect chamber contents to determine what's being stored and where it's located. | Remove product and allowing unit to recover. Review actions from the basic visuals above. | N/A | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
| Set point settings are incorrect. |  | Check the User Interface Temperature Setting for an accurate set point. Ensure the UI settings match setting in the downloaded diagnostics file. |  |  | No tools needed | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Inspect UI wiring harness for proper install on the power supply | Reseat the UI wiring harness on the power supply |  | Precision Tweezers for pin adjustment | Confirm UI display is functional. | No
|  |  | Ensure current firmware version is installed | Check UI Tools/Firmware for version | firmware BIN files | No tools needed | Verify firmware installed and Monitor pulldown | No
| An air sensor is dislodged or damaged. | Air1 or Air2 outside of range should be inspected further | Inspect location and ensure properly installed from factory setting. | Reinstall the air sensor into factory location. | Air Sensor | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Confirm temperature readings of air sensor using an external thermometer | Check calibration settings for improper offsets, replace sensor. | N/A | Fluke Thermometer, Calibration instructions from product manual | Power on unit, Monitor pulldown temperature based on pulldown chart | Sensor replacability to be confirmed
|  |  | Ensure air sensor cable is fully seated into the power supply | Inspect air sensor element or cable assembly for damage. Remove air sensor from clip inside of chamber; Remove air sensor cable assembly for damage; Disconnect air sensor harness from power supply. Reseat connector to power supply. |  | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
| Glycol sensor needs attention. | Glycol_temperature outside of range should be inspected further | Inspect glycol sensor & bottle to ensure properly filled with glycol, properly installed in harness, wiring is properly seated in the center area of the bottle. | Replace if needed |  |  | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Inspect glycol sensor cable ensuring jack is plugged into the glycol port and is fully seated into jack. | Reseat jack into port. If jack is damaged or part is remaining in the port, carefully remove the fragment with tweezers from the port and replace the glycol sensor |  | Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Ensure air & glycol cables are fully seated into the power supply | Disconnect air sensor harness from power supply. Reseat connector to power supply. |  | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Confirm temperature readings of glycol using an external thermometer | Replace if needed. | Glycol Bottle | Fluke Thermometer | Monitor pulldown temperature based on pulldown chart | No
| Chamber Fan Not Operating correctly | Need a line for chamber fan | While very unusual, inspect for ice buildup behind the fan in the duct. | First activate defrost operation from the UI. Thaw if needed, Identify how customer is using product that can cause ice to form inside the chamber. | N/A | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Ensure Fan power supply has power | Check voltage to power supply & fan connectors. Replace power supply if needed. | Internal Fan Kit | Phillips Screwdriver, Precision tweezers | Power on unit, confirm internal fan is operational | No
|  |  | Ensure internal fan is operating correctly and no debris is caught in the fan. | Replace if needed. | Internal Fan Kit | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
| Power Supply / Controller hardware failure |  | Confirm TEC firmware version | Replace power supply | Power Supply Kit | Phillips screwdriver | Power on unit, Monitor pulldown temperature based on pulldown chart | Yes due to cost and lead time
|  |  | Ensure the Power Supply. Is functioning properly. Was there a critical failure warning on the UI screen? Codes 0x0001 (faulty logic voltage), 0x0100 (voltage over) and 0x0200 voltage below critical requirements. Check the TEC firmware version to confirm that it is XX | Check the comm cable connestions between the UI and Power Supply control board. | N/A | Phillips screwdriver | Power on unit, critical error replaced by normal display screens. Monitor pulldown temperature based on pulldown chart | Yes if critical failure does not clear
|  |  | Power Supply Communication error shown on the UI screen | Reset the unit to reset system communications. If continues, check the power supply to UI cable to make sure it is properly installed. Inspect that the green LED is lit. | N/A | Phillips screwdriver | UI screen is functional. | No
|  |  | Inspect PCBA and all fans for dust or debris | Clean unit to remove dust & debris | N/A | Vacuum, Phillips screwdriver | Power on unit, Monitor pulldown temperature based on pulldown chart | No
| Temperature offsets inaccurate | Check Air Offset and Glycol Offset | Inspect the calibration offsets in the User Interface to ensure settings are not erratic. | Calibration instructions from product manual | N/A | N/A | Monitor pulldown temperature based on pulldown chart | No
Large Temperature Fluctuations
Large Temperature Fluctuations | Air sensor is dislodged or damaged. | Air1 or Air2 outside of range should be inspected further | Inspect location and ensure properly installed from factory setting. | Reinstall the air sensor into factory location. | Air Sensor | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Ensure air sensor cable is fully seated into the power supply | Inspect air sensor element or cable assembly for damage. Remove air sensor from clip inside of chamber; Remove air sensor cable assembly for damage; Disconnect air sensor harness from power supply. Reseat connector to power supply. |  | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | Sensor replacability to be confirmed
| Glycol sensor needs attention. | Glycol_temperature outside of range should be inspected further | Inspect glycol sensor & bottle to ensure properly filled with glycol, properly installed in harness, wiring is properly seated in the center area of the bottle. | Replace if needed |  |  | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Inspect glycol sensor cable ensuring jack is plugged into the glycol port and is fully seated into jack. | Reseat jack into port. If jack is damaged or part is remaining in the port, carefully remove the fragment with tweezers from the port and replace the glycol sensor |  | Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Ensure air & glycol cables are fully seated into the power supply | Disconnect air sensor harness from power supply. Reseat connector to power supply. |  | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Confirm temperature readings of glycol using an external thermometer | Replace if needed. | Glycol Bottle | Fluke Thermometer | Monitor pulldown temperature based on pulldown chart | No
| Chamber Fan Not Operating correctly | Need a line for chamber fan | While very unusual, inspect for ice buildup behind the fan in the duct. | First activate defrost operation from the UI. Thaw if needed, Identify how customer is using product that can cause ice to form inside the chamber. | N/A | N/A | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Ensure Fan power supply has power | Check voltage to power supply & fan connectors. Replace power supply if needed. | Internal Fan Kit | Phillips Screwdriver, Precision tweezers | Power on unit, confirm internal fan is operational | No
|  |  | Ensure internal fan is operating correctly and no debris is caught in the fan. | Replace if needed. | Internal Fan Kit | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
Glycol reading of less than -10C
Glycol reading of less than -10C | Glycol sensor needs attention. | Inspect Glycol_temperature and Glycol offset | Inspect glycol sensor & bottle to ensure properly filled with glycol, properly installed in harness, wiring is properly seated in the center area of the bottle. | Replace if needed |  |  | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Inspect glycol sensor cable ensuring jack is plugged into the glycol port and is fully seated into jack. | Reseat jack into port. If jack is damaged or part is remaining in the port, carefully remove the fragment with tweezers from the port and replace the glycol sensor |  | Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Ensure air & glycol cables are fully seated into the power supply | Disconnect air sensor harness from power supply. Reseat connector to power supply. |  | Phillips Screwdriver, Precision tweezers | Power on unit, Monitor pulldown temperature based on pulldown chart | No
|  |  | Confirm temperature readings of glycol using an external thermometer | Replace if needed. | Glycol Bottle | Fluke Thermometer | Monitor pulldown temperature based on pulldown chart | No
Door will not open
Door Will not open | Remains locked | Need a line for door lock status | Do you have power to the unit? If not is the battery dead? | Replace battery | CR123A |  | Can you unlock the door by pressing on the screen for 3 seconds to pull up the UI. | No
|  |  | Door lock harness | Inspect harness for damage or loose connection. | N/A | N/A | Inspect basic visual above for signs of cause of ice | No
|  |  | Lock catch | Push door in to eliminate tension. Push unlock |  |  |  | No
Shelves are not sturdy
Shelves are not sturdy | Check for damage to pilasters, clips, or shelf planarity | N/A | Are the clips in level? | Replace parts as needed. | Shelf Clip / Shelf (Optional) | N/A | Confirm level & secure shelving | No
UI screen is black when plugged in.
UI black screen | UI display is not powered on. | N/A | Is there a black screen on the display during power up? | Reset the unit to reboot the firmware | N/A | Phillips Screwdriver | Screen powers on as expected. | No
|  |  | Is screen is not lit up? Inspect cables to ensure properly seated. Ensure unit is plugged in and power supply is on. | Reseat cables, Reset unit, Then reseat cable into the power supply. |  |  | Power on unit, screen powers on as expected | No
UI screen is frozen
UI screen is frozen | Will not accept touch screen input | N/A | UI Screen | Does pressing the reset button fix the problem | N/A | N/A | Pressing the reset button restarts the refrigerator and the touch screen functionality is restored | No
|  |  |  | Replace the touchscreen | UI Board | Phillips Screwdriver | Restart the refrigerator and ensure that touch screen functionality is restored. | Yes due to cost and lead time
UI graphics are not displayed properly
UI graphics are not displayed properly | UI graphics are not properly displayed or visible | N/A | Attempt to reinstall firmware. | Attempt firmware reinstall |  | UI Firmware BIN file on USB memory device | Confirm firmware version on the UI | No
Excessive internal condensation
Internal Condensation | Air leak into the cabinet | N/A | Door seals, Passthrough port | Replace door seal, add putty to pass through port | Door seal | NA | Check air for any air leaks | No
| Forced air system malfunction |  | Internal fan, CO2 charge, PHP | Replace internal fan, Recharge system, Replace PHP | Internal fan, CO2 charging kit, PHP | Screwdriver, CO2 charging kit, Torque driver | Run for 24 hrs and check for reduced condensation. | Yes
Water pooling inside the cabinet
Water Pooling in the bottom of the cabinet | Clogged drain | N/A | Drain | Use a long drill bit to chase out the hole. Foam could be blocking the drain. Vacuum out any loose debris | None | Long drill bit | Place a small amount of water on the drain and make sure it is able to clear. | No
Water pooling on the floor below the unit
Water pooling under the unit | Evaporator malfunction. | N/A | Condensation tray | Reference service manual for proper evaporator installation. If installed properly, replace the condensation tray. | Condensation Tray | Screwdriver | Install tray without connecting thermistor. Check in 10 minutes to see if the tray is warm. If warm connect the thermistor and finish install | Maybe due to lead time
Cannot hear audible alarms
Sound not working | Cannot hear audible alarm or doesn't sound an audible alarm | N/A | UI Board | Replace UI Board | UI Board | Flathead Screwdriver | Confirm UI is functional, Test door ajar alarm | Yes, cost and lead time
Cannot set temperature
Cannot set temperature | Inaccurate temperature display | N/A | See steps in the temperature variability section | N/A | N/A | N/A | N/A
| Cannot set temperature | N/A | Ensure that the user has the correct privileges to change the set point | Login with the administrator. |  | N/A | Confirm UI is functional | No
USB drive not recognized
USB drive not recognized | USB not recognized or during data download process "download complete" screen doesn't display (indicates issues with USB drive). | N/A | Try a different device, does it work? | If not, replace UI Board | UI Board | N/A | Reattempt download function | No
| USB not recognized or during data download process "download complete" screen doesn't display (indicates issues with USB drive). | N/A | Wiring harness for the UI board to the power supply | Reseat wires if needed, Replace USB if needed. Replace grounding wires if damaged. | UI harness | Phillips Screwdriver | Reattempt download function | No
| USB not recognized | N/A | Ensure USB device is not locked or has available memory. Does the USB work on a PC? | Make sure that is it is formatted FAT32 or exFAT. See formatting instructions in service manual |  | USB memory stick | Reattempt download function | No
| USB not recognized or during data download process "download complete" screen doesn't display (indicates issues with USB drive). | N/A | Ensure no antivirus or encryption software is on the device. Does the USB work on a PC. | Make sure that is it is formatted FAT32 or exFAT and not write protected or unencrypted. See formatting instructions in service manual |  | USB memory stick | Reattempt download function and ensure download completes | No
Data not logging
Data not logging | Data not logging | N/A | Is the LED light solid Red ? Data log is full. | Download data & erase the data log through the UI firmware. See user manual for details. | N/A | USB memory stick | Download data and clear log, light will not change until the next data log interval. | No
Data not logging | Data not logging | N/A | Ensure firmware settings are set to log data | Review diagnostic file for confirmation of settings: Service Manual | N/A |  | Check diagnostic file downloaded during data log for accurate settings. | No
Cabinet lighting not working
Cabinet lighting not working | Improper Settings | N/A | When opening the door do all the lights fail to light up | Change the lighting settings to work when the door is opened | N/A | N/A | Open door, ensure that the light comes on | No
Cabinet lighting not working | USB Lighting Failure | N/A | Is only one light pod failing to work. | Replace the light pod | LED replacement | N/A | Open door, ensure that the light comes on | No
Red light blinking
Red light blinking | Red light blinking | N/A | Is the LED light blinking Red? No power to unit | Confirm that the unit is plugged into an active power source |  |  | Once power is reestablished download data and clear log | No

## System Alerts
Error Code | Error Detection | Description | First Step | 2nd Step / Replacement | System Error Message Resolution
0x0001 | error condition detected after 60 minutes of runtime; | Incorrect power/voltage for PHP1 | Remove AC power; Remove unit top cover; unplug PHP harness from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; Replace PHP; Install top cover; Apply AC power | error condition is immediately cleared;
0x0002 | error condition detected after 60 minutes of runtime; | Incorrect power/voltage for PHP2 | Remove AC power; Remove unit top cover; unplug PHP harness from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; Replace PHP; Install top cover; Apply AC power | error condition is immediately cleared;
0x0004 | error condition detected after 100 seconds of runtime; | Chamber fan reading out-of-bounds | Remove AC power; Remove unit top cover; unplug foamed-in harness connector from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; plug foamed-in harness connector into control board assembly; Install top cover; Apply AC power | error condition is immediately cleared;
0x0008 | error condition detected after 100 seconds of runtime; | Chamber 1 air sensor below limit | Remove AC power; Remove unit top cover; unplug foamed-in harness connector from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; plug foamed-in harness connector into control board assembly; Install top cover; Apply AC power | error condition is immediately cleared;
0x0010 | error condition detected after 100 seconds of runtime; | Chamber 2 air sensor below limit | Remove AC power; Remove unit top cover; unplug foamed-in harness connector from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; plug foamed-in harness connector into control board assembly; Install top cover; Apply AC power | error condition is immediately cleared;
0x0020 | after power on reset, all system action is halted (no power, fans, or communications) | Fan controller communications failure | HARDWARE MODIFICATION REQUIRED: Remove and disassemble power supply assembly; Modify control board to short onboard i2c sda and scl signals; Re-assemble unit; Apply AC power | HARDWARE MODIFICATION REQUIRED: Remove and disassemble power supply assembly; Modify control board to remove the shorting of the onboard i2c sda and scl signals; Re-assemble unit; Apply AC power | error condition does not appear after power-on reset;
0x0040 | error condition detected after 30 seconds of runtime | Reject sensor 1 or 2 reading below limit | Remove AC power; Remove unit top cover; unplug reject sensor harness connector from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; plug reject sensor harness connector into control board assembly; Install top cover; Apply AC power | error condition does not appear after power-on reset;
0x0080 | error condition detected after 30 seconds of being above 60C; Verity that power to cooling system is removed | Reject sensor 1 or 2 reading below limit | Remove AC power; Remove unit top cover; Apply AC power; Use proper precautions for high-voltage safety; Aim heat gun at any rear reject sensor until reject temperature exceeds 60C and system alert engages | Remove AC power; Remove unit top cover; Apply AC power; Use proper precautions for high-voltage safety; Aim heat gun at any rear reject sensor until reject temperature exceeds 60C and system alert engages. Remove heat gun from rear reject sensor | error condition is immediately cleared once reject temperature goes below 60C; verify that power to cooling system is restored.;
0x0100 | error condition detected after 100 seconds of runtime | Reject fan 1 or 2 out-of-bounds | Remove AC power; Remove unit top cover; unplug reject fan harness connector from control board assembly; Install top cover; Apply AC power | Remove AC power; Remove unit top cover; plug reject fan harness connector into control board assembly; Install top cover; Apply AC power | error condition is immediately cleared;

## Display Screen Shots
Status Message
Entry Indicators
Delete Button
Alarm Count Indicators
36

## Diagnostic File - Standard
File Output | Explanation
Date 02/22/2017, Time 11:30:52 | Date & Time data was downloaded
SN: XXXXXXXXXX1234 | Unit Serial Number
UIFirmware = TFB1A5 | Identifies firmware version
TECFirmware = F8 | Identified TEC (controller) firmware version. Should be CC
Setpoint_temp = 04.0C | UI set point temperature (Main Menu --> Settings --> Temp)
Control/Air temperature = 4.1C | Actual temperature of Air Sensor when downloaded
Control/Air offset = -0.9C | Calibration offsets programmed in for the Air sensor
Precision = DEGREES | Display setting of temperature resolution.
Upper Temp Threshold = 08.0C | UI upper temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Lower Temp Threshold = 02.0C | UI lower temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Glycol_temperature = 3.9C | Actual temperature of Glycol Sensor when downloaded
Glycol offset = 0.8C | Calibration offsets programmed in for the Glycol sensor
Displaying and Alerting on Glycol temp | What are the display settings that also trigger alarms.
Log_interval = 0015 | Defines frequency of data log - 15 represents 15 minutes. Options include 15, 60, 480 or 720
Datalog Memory alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Data_capacity = 0099 | Memory Remaining (%)
Data Threshold = 30.0 | What is the % remaining for a memory alarm to trigger
Battery Threshold = 30.0 | What is the % remaining for a battery alarm to trigger
Battery_Voltage = 02.8V | Represents the voltage of the battery. Value of 3 represents a fresh battery, 0 represents a dead / no battery
Battery_Capacity_Level = 0050 | Represents the percentage of battery capacity remaining. If below the Battery Threshold level then Battery Alarm would be active
Battery out of range in last minute count = 0000 | >0 means that you should replace the battery
Battery alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Door open seconds = 00000000 | If in a door ajar state what is the time the door has been open
Door opened count = 00000025 | How many times has the door been opened since the last data log reset
Door ajar audible alert seconds = 00000060 | What is the door ajar duration for an alarm to trigger (1 minute)
Door ajar audio alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
System error code = 0x0000 | Multiple Errors need to be added together
Power alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Defrost active = false | True indicates defrost cycle is active, False is not active
Has met initial pull down condition | True indicates temp range alarms are enabled, False means unit hasn't reached set point +3 C and temp range alarms are not engaged

## Diagnostic File - Glycol
File Output | Explanation | Troubleshooting guide line 60
Date 02/22/2017, Time 11:30:52 | Date & Time data was downloaded
SN: XXXXXXXXXX1234 | Unit Serial Number
UIFirmware = TFB1A5 | Identifies firmware version
TECFirmware = F8 | Identified TEC (controller) firmware version. Should be CC
Setpoint_temp = 04.0C | UI set point temperature (Main Menu --> Settings --> Temp)
Control/Air temperature = 4.1C | Actual temperature of Air Sensor when downloaded
Control/Air offset = -0.9C | Calibration offsets programmed in for the Air sensor
Precision = DEGREES | Display setting of temperature resolution.
Upper Temp Threshold = 08.0C | UI upper temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Lower Temp Threshold = 02.0C | UI lower temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Glycol_temperature = -57.0C | Actual temperature of Glycol Sensor when downloaded
Glycol offset = 0.8C | Calibration offsets programmed in for the Glycol sensor
Displaying and Alerting on Glycol temp | What are the display settings that also trigger alarms.
Log_interval = 0015 | Defines frequency of data log - 15 represents 15 minutes. Options include 15, 60, 480 or 720
Datalog Memory alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Data_capacity = 0099 | Memory Remaining (%)
Data Threshold = 30.0 | What is the % remaining for a memory alarm to trigger
Battery Threshold = 30.0 | What is the % remaining for a battery alarm to trigger
Battery_Voltage = 02.8V | Represents the voltage of the battery. Value of 3 represents a fresh battery, 0 represents a dead / no battery
Battery_Capacity_Level = 0050 | Represents the percentage of battery capacity remaining. If below the Battery Threshold level then Battery Alarm would be active
Battery out of range in last minute count = 0000 | >0 means that you should replace the battery
Battery alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Door open seconds = 00000000 | If in a door ajar state what is the time the door has been open
Door opened count = 00000025 | How many times has the door been opened since the last data log reset
Door ajar audible alert seconds = 00000060 | What is the door ajar duration for an alarm to trigger (1 minute)
Door ajar audio alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
System error code = 0x0000 | Multiple Errors need to be added together
Power alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Defrost active = false | True indicates defrost cycle is active, False is not active
Has met initial pull down condition | True indicates temp range alarms are enabled, False means unit hasn't reached set point +3 C and temp range alarms are not engaged

## Diagnostic File - Colder
File Output | Explanation | Troubleshooting Guide line 30
Date 02/22/2017, Time 11:30:52 | Date & Time data was downloaded
SN: XXXXXXXXXX1234 | Unit Serial Number
UIFirmware = TFB1A5 | Identifies firmware version
TECFirmware = F8 | Identified TEC (controller) firmware version. Should be CC
Setpoint_temp = 04.0C | UI set point temperature (Main Menu --> Settings --> Temp)
Control/Air temperature = 2.1C | Actual temperature of Air Sensor when downloaded
Control/Air offset = -0.9C | Calibration offsets programmed in for the Air sensor
Precision = DEGREES | Display setting of temperature resolution.
Upper Temp Threshold = 08.0C | UI upper temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Lower Temp Threshold = 02.0C | UI lower temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Glycol_temperature = 2.9C | Actual temperature of Glycol Sensor when downloaded
Glycol offset = 0.8C | Calibration offsets programmed in for the Glycol sensor
Displaying and Alerting on Glycol temp | What are the display settings that also trigger alarms.
Log_interval = 0015 | Defines frequency of data log - 15 represents 15 minutes. Options include 15, 60, 480 or 720
Datalog Memory alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Data_capacity = 0099 | Memory Remaining (%)
Data Threshold = 30.0 | What is the % remaining for a memory alarm to trigger
Battery Threshold = 30.0 | What is the % remaining for a battery alarm to trigger
Battery_Voltage = 02.8V | Represents the voltage of the battery. Value of 3 represents a fresh battery, 0 represents a dead / no battery
Battery_Capacity_Level = 0050 | Represents the percentage of battery capacity remaining. If below the Battery Threshold level then Battery Alarm would be active
Battery out of range in last minute count = 0000 | >0 means that you should replace the battery
Battery alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Door open seconds = 00000000 | If in a door ajar state what is the time the door has been open
Door opened count = 00000025 | How many times has the door been opened since the last data log reset
Door ajar audible alert seconds = 00000060 | What is the door ajar duration for an alarm to trigger (1 minute)
Door ajar audio alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
System error code = 0x0000 | Multiple Errors need to be added together
Power alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Defrost active = false | True indicates defrost cycle is active, False is not active
Has met initial pull down condition | True indicates temp range alarms are enabled, False means unit hasn't reached set point +3 C and temp range alarms are not engaged

## Diagnostic File - Critical
File Output | Explanation | Troubleshooting guide line #9
Date 02/22/2017, Time 11:30:52 | Date & Time data was downloaded
SN: XXXXXXXXXX1234 | Unit Serial Number
UIFirmware = TFB1A5 | Identifies firmware version
TECFirmware = F8 | Identified TEC (controller) firmware version. Should be CC
Setpoint_temp = 04.0C | UI set point temperature (Main Menu --> Settings --> Temp)
Control/Air temperature = 7.1C | Actual temperature of Air Sensor when downloaded
Control/Air offset = -0.9C | Calibration offsets programmed in for the Air sensor
Precision = DEGREES | Display setting of temperature resolution.
Upper Temp Threshold = 08.0C | UI upper temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Lower Temp Threshold = 02.0C | UI lower temp threshold for Temp Alarm (Main Menu --> Alarms --> Temp Range)
Glycol_temperature = 6.9C | Actual temperature of Glycol Sensor when downloaded
Glycol offset = 0.8C | Calibration offsets programmed in for the Glycol sensor
Displaying and Alerting on Glycol temp | What are the display settings that also trigger alarms.
Log_interval = 0015 | Defines frequency of data log - 15 represents 15 minutes. Options include 15, 60, 480 or 720
Datalog Memory alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Data_capacity = 0099 | Memory Remaining (%)
Data Threshold = 30.0 | What is the % remaining for a memory alarm to trigger
Battery Threshold = 30.0 | What is the % remaining for a battery alarm to trigger
Battery_Voltage = 02.8V | Represents the voltage of the battery. Value of 3 represents a fresh battery, 0 represents a dead / no battery
Battery_Capacity_Level = 0050 | Represents the percentage of battery capacity remaining. If below the Battery Threshold level then Battery Alarm would be active
Battery out of range in last minute count = 0000 | >0 means that you should replace the battery
Battery alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Door open seconds = 00000000 | If in a door ajar state what is the time the door has been open
Door opened count = 00000025 | How many times has the door been opened since the last data log reset
Door ajar audible alert seconds = 00000060 | What is the door ajar duration for an alarm to trigger (1 minute)
Door ajar audio alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
System error code = 0x0080 | Multiple Errors need to be added together
Power alarm enabled = true | True indicates door ajar alarm enabled, False is disabled.
Defrost active = false | True indicates defrost cycle is active, False is not active
Has met initial pull down condition | True indicates temp range alarms are enabled, False means unit hasn't reached set point +3 C and temp range alarms are not engaged